## Slide 1: Absolute Charge Measurement in UHDpulse electron beams
F. Frei, A. Schüller, C.Bailat, P. Peier

## Slide 2: Calibration of absolute charge measurement
Reference charge pulse generator (PTB)
A1.1.3

## Slide 3: Calibration of absolute charge measurement
Reference charge pulse generator (PTB)
ICTs and ACCT@METAS

## Slide 4: Calibration of absolute charge measurement
ICT signal
filtering
26dB gain
raw signal from accelerator
Verifications
Using accelerator:
Connecting full measurement chain including proper readout electronics
Record traces with respect to timing windows
Check center of mass
Check relative signal contributions in timing windows
Connecting calibrated reference charge pulse generator:
Set proper pulse length (3us)
Adjust delay to match situation with accelerator (center of mass with respect to timing windows)
Verify relative signal contributions in timing windows
Baseline
+
-

## Slide 5: Calibration of absolute charge measurement
With calibrated reference charge pulse generator:
Expanded uncertainty of deduced sensitivity: 1% (k=2) Dominant contribution from: correction factor of nominal pulse charge given as a function of pulse length for each current setting.
26dB gain
32dB gain

## Slide 6: Absolute charge measurement used for Fricke dosimetry
number of electrons
Fricke total absorption:
Fricke dosimeter in bags:
Comparison of the primary standards of PTB and METAS for absorbed dose to water in ultra-high dose per pulse electron beams via alanine dosimeters provided by the National Research Council of Canada (NRC).
The summary of this analysis is that the standards of the METAS and PTB for UHDR electron beams agree within the combined standard uncertainty of the comparison measurement, RMETAS/PTB = 1.002 ± 0.012 (k=1). -> Deliverable 3 (See talk 24. of Alexandra Bourgoin, PTB).

## Slide 7: Calibration of ACCT
Chart: Current setting reference charge pulse generator [mA]
| 100 | 30 | 15 | 10 | 5 | 3 | 1.5 | 1 | 0.5 | 0.3
 | 99.143 | 30.087 | 15.067 | 9.969 | 4.973 | 2.993 | 1.508 | 1 | 0.499 | 0.305
CHUV ACCT
METS ACCT
@METAS:
Existing calibration in good agreement with the calibrated reference charge pulse generator.
Since the ACCT signal is only used for rough beam setup, the calibration coefficient were not updated.
accelerator

## Slide 9: Thank you very much for your attention