## Slide 1: Primary standard calorimetry for UHDR proton beams
Anna Subiel, PhD, CPhys
Medical Radiation Science, National Physical Laboratory, Hampton Road, Teddington, Middlesex, TW11 0LW, UK
UHDpulse 2nd Stakeholder Meeting, 26-27 January 2023

## Slide 2: Calorimeters used as prim. standards in high energy beams
D=c·ΔT
Vm
Vout
R4
R3
R2
Rth+wire
0.0E-07
5.0E-07
1.0E-06
1.5E-06
2.0E-06
2.5E-06
3.0E-06
-200
-100
0
100
200
Time (s)
V
out
(V)
ΔT
thermistor
calorimeter phantom
isolation
bridge circuit
bridge output signal

## Slide 3: Calibration in high energy beams
For high-energy photon & electron beams:
calibrations based on a primary standard calorimeters for the direct determination of absorbed dose to water with uncertainties on reference dosimetry of:
0.65% (k=1) for photons
0.75% (k=1) for electrons
X-rays
NPL electron calorimeter
NPL photon calorimeter

## Slide 4: Calibration in proton beams

## Slide 5: World’s 1st Primary-Standard Calorimeter for Protons
Graphite calorimeter
Originally developed for use with conventional proton beams,
Consists of graphite discs arranged in a nested construction, maintained under vacuum

## Slide 6: World’s 1st Primary-Standard Calorimeter for Protons
Graphite calorimeter
Originally developed for use with conventional proton beams,
Consists of graphite discs arranged in a nested construction, maintained under vacuum

## Slide 7: World’s 1st Primary-Standard Calorimeter for Protons
16 mm
White tiny spot - thermistor
Graphite calorimeter
4 thermistors – 0.4 mm diameter typical fraction of treatment - 2 Gy, 0.002 degrees
Delivers an uncertainty on reference dosimetry for protons <1% (k=1)

## Slide 8: Proton reference dosimetry
2000 – IAEA TRS 398
  Ionisation chambers
Uncertainty on for protons: 2.3% (k=1)
Detectors currently used in the proton clinics are calibrated in a Cobalt-60 beam
Dose in a proton beam
Conversion factor,
Analytical
Uncertainty in dose at least 2 times larger than the recommended uncertainty
Courtesy of A. Lourenço

## Slide 9: Proton reference dosimetry
Detectors currently used in the proton clinics are calibrated in a Cobalt-60 beam
Dose in a proton beam
Conversion factor,
Analytical
Calibrate the detectors directly in the same or similar beam used clinically
Courtesy of A. Lourenço

## Slide 10: Analogy between FLASH RT and proton RT

## Slide 11: Calorimetry in UHDR beams
NPL primary standard graphite calorimeter
the control and analysis software was reconfigured to enable it to be used with UHDR particle beams
Operated in quasi-adiabatic mode, thermistors detect changes in temperature of the graphite created by energy absorbed from the radiation beam allowing derivation of absorbed dose
Fig. NPL’s primary standard graphite calorimeter.

## Slide 14: Calorimetry in UHDR beams
NPL primary standard graphite calorimeter
250 MeV (Varian ProBeam® operating in research mode) at ~63 Gy/s, proton bunches with 0.2-ns pulse duration at 72.8-MHz RF repetition rate corresponding to 0.2-ns micro-pulses separated by 13.7-ns intervals
The core of the calorimeter was positioned at the isocentre with graphite plates placed in front to position the core at a WET of 5.2 gcm-2
    5×6 cm2
    5×8 cm2
    5×10 cm2
    5×12 cm2
    6×5 cm2
    12×5 cm2
Experimental setup mounted on the patient couch and radiograph of the NPL PSPC. (a) (1) Gantry, (2) NPL PSPC, (3) measurement instrumentation for the NPL PSPC, (4) vacuum pump, (5) ionisation chamber setup.

## Slide 15: Calorimetry in UHDR beams
Data were collected at 5 Hz and irradiations were repeated at least 20 times for each field tested to achieve a relative standard uncertainty of 0.04% (Type A)
The average measured absorbed dose for the six fields by the NPL PSPC was 7.8 Gy with a standard deviation of 0.5%.
Beam
direction
Lourenço et al. Sci.Rep. accepted
Calorimetry run in 250 MeV FLASH proton beam for 5x6 cm2
Notes: Irradiations were repeated every two minutes. For each irradiation, a quadratic fit as a function of time is applied to the pre- and post- irradiation drift curves. These drifts are extrapolated and the increase in temperature is determined from the offset between the two drifts (Fig. 1). The average absorbed dose to the core is determined by multiplying the measured increase in temperature by the specific heat capacity of the core, which was determined experimentally at NPL [28].
The absorbed dose-to-core is converted to absorbed dose-to-water by the product of the necessary beam-dependent correction factors which were determined using Monte Carlo simulations. (i) The impurity correction factor, 𝑘_imp, accounts for the presence of non-graphite components in the calorimeter geometry, such as thermistors, epoxy resin (to hold the thermistors), and thermistor wires. 𝑘_imp is determined as the ratio between the dose averaged over the core in a pure graphite geometry (with vacuum gaps) and the dose averaged over the core in the real geometry with all the non-graphite components. 𝑘_gap, accounts for the presence of vacuum gaps within the calorimeter to minimise heat transfers. 𝑘_gap was determined by modelling the calorimeter with a compensated geometry. The product 𝑘_imp. 𝑘_gap converts the dose-to-core in the real calorimeter geometry to dose-to-graphite in a geometry composed of homogeneous graphite. Dose-to graphite is converted to dose-to-water by (iii) the water-to-graphite mass-stopping-power ratio, 𝑠_(w,g), and (iv) the fluence correction factor, 𝑘_fl, to correct for the difference in fluence at water-equivalent depths between water and graphite. The 𝑠_(w,g) ratio is determined as an integral over the charged particle spectrum and 𝑘_fl determined from the charged particle fluence differential in energy between water and graphite [25].
𝑘_(z, cal) (corrects for the reference distance from the beam’s source).

## Slide 16: Calorimetry in UHDR proton beams
 First ever calorimetry measurements in UHDR proton beam
Established the correction factors required for absolute dosimetry of FLASH proton beam radiotherapy (Lourenço et al., 2023 Nature, Scientific Reports, accepted)  see also Deliverable 2 of the UHDpulse project
Measurement uncertainty of 0.9% (k=1) – in line with clinical requirement
Underpinned the FDA approval and provided the hospital with confidence to commence clinical implementation of this novel technology
NPL and Cincinnati Teams

## Slide 17: Thank you
anna.subiel@npl.co.uk
This project (18HLT04) has received funding from the EMPIR programme co-financed by the Participating States and from the European Union’s Horizon 2020 research and innovation programme.
http://uhdpulse-empir.eu/
Acknowledgements
Ana Lourenco
Nigel Lee
Hugo Palmans
Russell Thomas