## Slide 1: flashDiamond – Review and Outlook
R. Kranzer, M. Marinelli, G. Verona Rinati, A. Schüller, HK. Looe, B. Poppe
2nd Stakeholder Meeting January 2023

## Slide 2: flashDiamond – Starting point
Challenge of ultra-high dose rates (UHDR) > 40 Gy/s as investigated for FLASH-RT
Reasoning in the UHDpulse introduction paper (Schüller et al. 2020 DOI:10.1016/j.ejmp.2020.09.020)
Key features
  Suitable for clinical electron beam dosimetry in conventional RT and IORT
  Radiation hardness with accumulated dose

## Slide 3: flashDiamond – Fundamental investigations
Investigations of the microDiamond presented in the following papers
    Marinelli et al. 2022 (DOI: 10.1002/mp.15473)
    Kranzer et al. 2022 (DOI: 10.1088/1361-6560/ac594e)
Key findings
  The microDiamond is not generally suitable for UH-DPP
  The crucial properties with respect to the linearity are
    Sensitivity
    Doping concentration in the p-type layer of the diamond Schottky-diode respectively the total series resistance
The optimization of these parameters leads to a novel type of detector, the flashDiamond

## Slide 4: flashDiamond – Fundamental investigations
The applicability for the commissioning of FLASH RT electron beams was shown in the following work
  G. Verona Rinati et al. 2022 (DOI: 10.1002/mp.15782)
Successfully tested in terms of:
  Linearity with ultra-high dose per pulse (UH-DPP)
  Depth dose curves (PDD)
  Lateral dose profiles
  Output factors
So far, all measurements were done at the PTB research accelerator or a SIT machines

## Slide 5: flashDiamond – Varian eFLASH
Upgrade of Clinac systems (eFLEX)
Beam parameters
  Beam quality 16 MeV
  Pulse duration 4.2 µs
  Pulse repetition frequency 180 Hz
Linearity up to 6 Gy per pulse

## Slide 6: flashDiamond – Varian eFLASH
Depth dose curves
  Conventional mode
    Comparison against Advanced Markus chamber (dose converted)
  FLASH mode
    Comparison against Advanced Markus chamber (dose converted and corrected for recombination)
flashDiamond enables PDD measurement in CONV and FLASH mode without dose conversion and corrections
Differences in the bremsstrahlung background between CONV and FLASH can be attributed to different scattering foils

## Slide 7: flashDiamond – IntraOp FLASH Mobetron
Beam parameters
  Beam quality 9 MeV
  Pulse duration 0.5 to 4 µs
  Pulse repetition frequency 30 Hz
  SSD 50 cm
  Field size 10 cm diameter
Linearity with varying pulse duration
  Pure using the nominal pulse duration
  Good using the “real” pulse duration
  Further improved using FWHM
Precise knowledge of the pulse duration is required for the determination of the linearity
Tools with high temporal resolution needed

## Slide 8: flashDiamond – IntraOp FLASH Mobetron
Beam parameters
  Beam quality 9 MeV
  Pulse duration 1 µs
  Pulse repetition frequency 15 Hz
  SSD 18.4 cm (max. 3 Gy per pulse)
  Field size 6 cm diameter
Depth dose curves
Lateral dose profiles
flashDiamond enables measurements of 3D dose distribution in CONV and FLASH mode without dose conversion and corrections

## Slide 9: flashDiamond - Applications
Relative Dosimetry
  Measurement of three-dimensional dose distribution with high spatial resolution
  Direct measurement of absorbed dose to water, no need to convert depth dose curves
Consistency check (beam stability)
Time resolved measurements (see talk of G. Verona Rinati)
Absolute dose measurements
  After cross-calibration against a suitable reference
  A methodology for the determination of absorbed dose to water is still missing

## Slide 10: flashDiamond - Proposal
Secondary standard for absorbed dose to water determination in conventional RT are ion chambers
Absorbed dose to water determination with diamond detectors?
        =

## Slide 11: flashDiamond - Proposal
Absorbed dose to water determination with diamond detectors
Beam quality correction factors for electron beams

## Slide 12: flashDiamond - Conclusion
Usability and properties validated
  on four different UHDPP electron accelerators types
  three of them are commercially available for research in FLASH RT
Methodology for the determination of absorbed dose to water
  Should be studied in more detail
  Determination of the beam quality correction factors and its uncertainty under reference conditions is in progress
The flashDiamond Type 60025 is commerciallized as a laboratory device for research purposes and thus available for further users and a variety of dosimetric applications

## Slide 13: Thank you!
Questions? Wishes? Notes? …