## Slide 1
VOLUME RECOMBINATION IN PARALLEL-PLATE IONIZATION CHAMBERS F. Gómez1,2, A. Schüller3, J. Paz-Martín1, A. Bourgouin3, R. Kranzer4,5, J. Pardo-Montero6, D. Gonzalez-Castaño2, N. Gómez-Fernández2, G. Felici7, D. Velleren8,9, A. Gasparini8,9,10, V. Vanreusel8,9,10
1 Universidad de Santiago, Santiago 15782, Spain;
2 Laboratorio de Radiofísica, Universidad de Santiago, Santiago 15782, Spain;
3 Physikalisch-Technische Bundesanstalt, Braunschweig 38116, Germany;
4 PTW-Freiburg (R&D), Freiburg 79115, Germany;
5 University Clinic for Medical Radiation Physics, Medical Campus Pius Hospital, Carl von Ossietzky University Oldenburg, 26121, Germany;
6 Group of Medical Physics and Biomathematics, Instituto de Investigación Sanitaria de Santiago (IDIS), Travesía da Choupana, Santiago de Compostela, A Coruña, 15706, Spain;
7 SIT, Aprilia, Italy;
8 Department of Radiotherapy, Iridium Network, Belgium;
9 Faculty of Medicine and Health Sciences, University of Antwerp, Belgium;
10 SCK CEN, Research in Dosimetric Applications, Mol, Belgium;
This work has received funding from the EMPIR programme project 18HLT04 UHDpulse co-financed by the Participating States and from the European Union’s Horizon 2020 research and innovation programme .

## Slide 2
The problem of recombination
UHDR and the recombination problem in PPIC
  Equations of charge carrier transport and numerical solutions
  Electron mobility and electron attachment
  Free electron fraction
  Contribution of electric field perturbation and electron-ion recombination
  Magnitudes of influence
  Pulse structure
  Instantaneous electrical current data
The importance of the electrode distance: Ultra Thin IC
  Saturation effect
  Charge multiplication
  Dose rate dependence
Some invariants from experimental data

## Slide 3
The problem of recombination
Dose per pulse (Gy)
CCE
Electron beam at MELAF (PTB)
Chamber PPC40
Pulse duration 2.5 µs
Frequency 5 Hz
Parallel plate ionization chamber in the transversal coordinate
x
x
0
d
E
n-
n+
x
0
d
E
n-
n+
x
0
d
E
n-
n+
Very Low DPP
Linear behaviour of carrier densities
Almost constant electric field
Negligible recombination
Low & Moderate DPP
Small non linear behaviour of carrier densities
Small screening of electric field
Aproximate analytical solutions for recombination
High & Ultra High DPP
Dominant non linear behaviour of carrier densities
Very large screening of electric field
No analytical solutions for recombination

## Slide 4
The problem of recombination
Ion mobility μ+(E) μ-(E)
Electron mobility ve(E)
Electron atachment τe(E)
Ion recombination rate α
Ionization
Charge Carrier densities
Recombination and attachment
Radiation model
Electric field
Charge transport
Electrical current
Deattachment/e-ion recombination / mobility /diffusion

## Slide 5
Equations of charge carrier transport
, are the positive, negative ion and electron densities
number of electron-ion pairs roduced by radiation escaping initial recombination per unit time and unit volume
is the recombination coefficient for ionic species, is the electron-ion recombination coefficient
, , are the charge carrier mobilities
is the electron attachment constant
In the ultra-high dose per pulse limit the air ionization chambers are subject to large recombination losses and substantial ksat correction factor values. Most approaches until now consider the one-dimensional equivalent problem:
Paz-Martin J. et al. Physica Medica 103, pp. 147-156 (2022)

## Slide 6
Equations of charge carrier transport
Numerical solutions are built through a discretization of time and coordinate between electrodes:
Paz-Martin J. et al. Physica Medica 103, pp. 147-156 (2022)
typically
s
N
Example of the accuracy of the simulation of a 2 mm chamber.
First versión of PPIC simulator released in April 2021

## Slide 7
Electron mobility
Paz-Martin J. et al. Physica Medica 103, pp. 147-156 (2022)
Electron drift velocity as a function of electric field is substantially different in humid air respect to dry air for moderate electric field ( ~100 V/mm)

## Slide 8
Electron attachment
Paz-Martin J. et al. Physica Medica 103, pp. 147-156 (2022)
Advanced
Markus
@ 300 V
UTIC
0.25 mm
@ 300 V
…
…
Three-body attachment
Two-body attachment
The use of very high electric can enhance the two-body attachment processes and yield an effective reduction of electron lifetime.

## Slide 9
The problem of recombination
Electron-ion recombination typical contribution to CCE is bellow 1%
(m3 s-1)
Models without electric field perturbation can not reproduce the problem in an accurate way.

## Slide 10
Free electron fraction
1) FEF Boag definition:
  Fraction of electrons liberated in the chamber that reach the positive electrode.
2) FEF signal:
  Fraction of the total charge readout of the chamber due to free electrons respect to saturation charge (neglecting volume recombination).
3) Effective FEF signal:
  Fraction of the total charge readout of the chamber respect to the total collected charge.
  This last quantity is the only one directly measurable!

## Slide 11
Free electron fraction (Roos chamber)
Values obtained from numerical simulation.
Voltage polarization 300 V and pulse duration 4.5 µs.
FEF Boag
FEF ind
FEF’ ind

## Slide 12
Free electron fraction (Roos chamber)
Experimental and simulated effective FEF’ind obtained from integrating the electrical current during the pulse.
“Numerical modeling of air-vented parallel plate ionization chambers for ultra-high dose rate applications” Phys Med. 2022 Oct 28;103:147-156. doi: 10.1016/j.ejmp.2022.10.006.
Effective free electron fraction for a 2.5 µs pulse in MELAF (PTB).

## Slide 13
Instantaneous current (Roos chamber)
“Numerical modeling of air-vented parallel plate ionization chambers for ultra-high dose rate applications” Phys Med. 2022 Oct 28;103:147-156. doi: 10.1016/j.ejmp.2022.10.006.
Instantaneous current for a 2.5 µs pulse in MELAF (PTB)

## Slide 14
Charge collection efficiency (Roos chamber)
“Numerical modeling of air-vented parallel plate ionization chambers for ultra-high dose rate applications” Phys Med. 2022 Oct 28;103:147-156. doi: 10.1016/j.ejmp.2022.10.006.
Chambers operated at 300 V for a pulse duration of 2.5 µs

## Slide 15
Charge collection efficiency (Adv Markus)
Agreement is not so perfect in general once volume recombination coefficient is fixed and treat the whole set of data equally.
Data from MELAF PTB for
2.5 s pulse duration and
m3 s-1
N.B.: Unpublished data.

## Slide 16
Pulse time structure (Roos chamber)
“Numerical modeling of air-vented parallel plate ionization chambers for ultra-high dose rate applications” Phys Med. 2022 Oct 28;103:147-156. doi: 10.1016/j.ejmp.2022.10.006.
CCE dependence on pulse duration for a 2 mm PPIC operated at 300 V bias voltage for a dose per pulse of 0.5 Gy.

## Slide 17: Charge collection efficiency and influence magnitudes
Typical sensitivity factors for CCE Advanced Markus (@ 1 Gy DPP and 1 μs)
Magnitude | Sensitivity
Temperature | 0.12 % K-1
Pressure | 0.03 % hPa-1
Voltage | 0.08 % V-1
Chamber gap | 0.06 % μm-1

## Slide 18: Charge collection efficiency and influence magnitudes
CCE Advanced Markus (pulse duration 2.5 μs)

## Slide 19: The importance of electrode distance
Gómez, F, Gonzalez-Castaño, DM, Fernández, NG, Pardo-Montero, J, Schüller, A, Gasparini, A, Vanreusel, V, Verellen, D, Felici, G, Kranzer, R, Paz-Martín, J. Development of an ultra-thin parallel plate ionization chamber for dosimetry in FLASH radiotherapy. Med Phys. 2022; 49: 4705– 4714.
https://doi.org/10.1002/mp.15668
Charge collection efficiency 99% @ 300 V

## Slide 20: The importance of electrode distance: Ultra Thin IC
“Charge collection efficiency, underlying recombination mechanisms, and the role of electrode distance of vented ionization chambers under ultra-high dose-per-pulse conditions”
Phys Med. 2022 Nov 7;104:10-17.
doi: 10.1016/j.ejmp.2022.10.021.
Charge collection efficiency of the PPICs with different electrode distances and three different chamber voltages. From left to right d = 1 mm, 0.5 mm, 0.25 mm. Experimental and simulated data as well as comparative data from the literature are presented

## Slide 21: The importance of electrode distance: UTIC
Measurement of UTIC charge vs. dose per pulse at water phantom PTB, pulse duration 2.5 µs.
Gómez, F, Gonzalez-Castaño, DM, Fernández, NG, Pardo-Montero, J, Schüller, A, Gasparini, A, Vanreusel, V, Verellen, D, Felici, G, Kranzer, R, Paz-Martín, J. Development of an ultra-thin parallel plate ionization chamber for dosimetry in FLASH radiotherapy. Med Phys. 2022; 49: 4705– 4714.
https://doi.org/10.1002/mp.15668
250 V

## Slide 22: The importance of electrode distance: UTIC
“Charge collection efficiency, underlying recombination mechanisms, and the role of electrode distance of vented ionization chambers under ultra-high dose-per-pulse conditions”
Phys Med. 2022 Nov 7;104:10-17.
doi: 10.1016/j.ejmp.2022.10.021.
Measured dose per pulse vs reference dose for an UTIC operated at 250 V in MELAF 20 MeV electron beam
pulse duration 2.5 µs.

## Slide 23: Charge multiplication
“Charge collection efficiency, underlying recombination mechanisms, and the role of electrode distance of vented ionization chambers under ultra-high dose-per-pulse conditions”
Phys Med. 2022 Nov 7;104:10-17.
doi: 10.1016/j.ejmp.2022.10.021.
Relative signal excess for UTIC at 500 V around 2 %.
First Townsend coefficient in air

## Slide 24: UTIC dose rate dependence
E-poster “ION RECOMBINATION DOSE RATE DEPENDENCE OF ULTRA THIN AIR IONIZATION CHAMBER” FRPT2022
CCE dependence on instantaneous dose rate is more evident for chambers with ion collection time similar to pulse duration.
CCE as a function of dose per pulse for different instantaneous dose rates for the Ultra-Thin Ionization Chamber (0.25 mm) operated at different polarization voltages: A) 100 V, B) 200 V and C) 300 V. Experimental data from MELAF (PTB) is shown with circles with uncertainty bars while numerical simulation predictions are indicated with solid lines.

## Slide 25: Some invariants from experimental data
Ion collection time defined at 99.95% of total collected charge for a Roos chamber show a linear relationship with inverse of the operating chamber voltage independently of the DPP (0.2 Gy to 5.8 Gy).

## Slide 26: Some invariants from experimental data
Under equal irradiation conditions CCE between two PPIC chambers is invariant if
“Charge collection efficiency, underlying recombination mechanisms, and the role of electrode distance of vented ionization chambers under ultra-high dose-per-pulse conditions”
Phys Med. 2022 Nov 7;104:10-17.
doi: 10.1016/j.ejmp.2022.10.021.

## Slide 27: Conclusions
Ultra Thin Air Ionization Chambers are able to work at DPP of 5.5 Gy with CCE > 99 % for 2.5 µs pulse duration and Vbias 250 V.
Pulse duration has some effect in chamber readout (ion collection time/delivery time). For UTIC it has very limited effect at nominal voltage.
Invariance of some quantities found in simulation and in data have to be properly explained from Physics.
Universal correction of large saturation effect is difficult due to the unit to unit variation, dependence on pulse duration (time structure) and effect of magnitudes of influence.

## Slide 28
Complexity vs model
“Non-equilibrium Air Plasmas at atmospheric Pressure” 2005 Institute of Physics, Ed. K H Becker, U. Kogelschatz, K H Schoenbach, R J Barker