## Slide 1: Ionisation Chamber Dosimetry in Ultra-Short Pulsed High Dose-Rate Very High Energy Electron (VHEE) Beams
Anna Subiel, PhD
Medical Radiation Science, National Physical Laboratory, Hampton Road, Teddington, Middlesex, TW11 0LW, UK
UHDpulse 2nd Stakeholder Meeting, 26-27 January 2023

## Slide 2
Most of the studies performed using electron beams accelerated by modified clinical LINAC or dedicated electron accelerators (E < 20 MeV)
DRAWBACKS
Limitations of clinical electrons
High relative surface dose
Shallow penetration/short range
Range straggling (no Bragg peak)
High penumbra
Bulging effect
Spread of beam in air (why we have cones)
FLASH RT
Radiation Oncology Physics, IAEA 2005
Limitations of electron beams due to energy – what happens if the electron energy is increased??
BEAM ENERGY

## Slide 3: What happens if the electron energy is increased (100+ MeV)?
Increases depth of penetration
No range straggling if beam penetrates through patient
Ability to control position electromagnetically
  Scanning beams more easily done than heavy particles
  Speed of electromagnetic scanning allows for ~ 100X more beams delivered in the same time as photons
Lower beam spread and reduced bulging effect
While maintaining some low energy characteristics
Can produce pencil beams
Less costly than heavy particles
High surface dose

## Slide 4
DesRosiers et al., Proceedings of the SPIE, Vol. 6881, (2008)
VHEE unlike X-ray photons maintain electronic equilibrium in tissues with varying densities
Perturbation of dose, - consequence - under or over dosage of tissue
Advantages of VHEE RT
conventional
RT photons
VHEE
Credit: C. DesRosiers

## Slide 5
Initial Recombination
Recombination along a single charged particle track.
Independent of dose and dose-rate.
More pronounced in highly ionising particles such as alpha-particles.
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
initial recomb.
general recomb.
Ionization chamber
General Recombination
Recombination between separate charged particle tracks.
Directly dependent on charge density i.e the number of ions produced per unit volume.
Dose-rate dependent.
General recombination is likely to play a much larger role in recombination effect in UHDR pulsed beams

## Slide 6
Dose-per-pulse [Gy]
Ion collection efficiency
IC: Metrological challenges of dosimetry at UHDR
Petersson et al., Med Phys 44 (2017) 1157
Typical behaviour of an IC at UHDR dose rates for pulsed electron beams
PTW Advanced Markus
(1 mm electrode separation)
ion recombination effect
6 MeV e-beam
ion collection time: 22 us
no physical meaning

## Slide 7
First experimental results: UHPDR VHEEs
ion collection time: 125 us
PTW Roos chamber
(2 mm electrode separation)
Notes: The calorimeter and chamber were exposed to a quasi-monoenergetic pulsed electron beam of approximately 200 MeV at the CLEAR user facility at CERN [3]. Each electron pulse was formed of a variable number of shorter electron bunches, with an adjustable charge-per-bunch range between 0.001 nC to 1.5 nC. Each bunch had a length of approximately 1 ps. By adjusting the charge-per-bunch and the number of bunches-per-pulse, we investigated charge-per-pulse values ranging from 0.05 nC to 11 nC. The collection time of the ion chamber is of the order of μs, therefore the chamber is sensitive to the total macro pulse and not the individual bunches.
The absolute value of the recombination factor was calculated from the ratio of the absolute dose determined from the calorimeter and the dose calculated from the ion chamber: 𝑘_(𝑠,𝑎𝑏𝑠)=𝐷_(𝑤,𝑐𝑎𝑙)/(𝑀𝑘_𝑝𝑜𝑙 𝑘_𝑇𝑃 〖𝑘_(𝑄,𝑄_0 ) 𝑘_(𝑄,𝑄_0 ) 𝑁〗_(𝐷,𝑤,𝑄_0 ) )
where 𝐷_(𝑤,𝑐𝑎𝑙) is the dose-to-water measured from the calorimeter, M is the charge reading of the chamber, 𝑘_𝑝𝑜𝑙 is the polarity correction factor, 𝑘_𝑇𝑃 is the temperature and pressure correction factor, 𝑘_(𝑄,𝑄_0 )is the beam quality correction factor and 𝑁_(𝐷,𝑤,𝑄_0 ) is the calibration coefficient of the ionisation chamber in the reference beam quality
The dose-to-water conversion factor, used to convert from the measured dose-to-graphite from the calorimeter into the required dose-to-water, was calculated to be 1.0912. The dose-per-pulse to water and the recombination factors determined from absolute measurements are displayed in Table 1.

## Slide 8
Experimental results: UHPDR VHEEs
OBJECTIVE: To study ion collection efficiency as a function of dose-per-pulse at instantaneous dose rates 5.0 × 106–3.1 × 108 Gy/s for VHEE beams
BEAM PARAMETERS: 200 MeV, x and y 𝜎 of 5 mm, ΔE between 0.25 and 6.5%
side-by-side measurements: PTW Roos chamber and NPL’s graphite calorimeter operating in quasi-adiabatic mode, where and = 1.103 (±0.063%); =0.009% (±0.009%) for k=1
quantification of the recombination factor ks,abs for the Roos chamber for a range of collecting voltages: 75 V – 600 V
BEAM DIRECTION
MONITOR CHAMBER
CALORIMETER
IC
The test-stand at the CLEAR facility, with the calorimeter, ion chamber and monitor chamber placed along the beam line with the beam travelling from right to left.
Nominal Beam Charge (nC/pulse) | Dw,cal (Gy/pulse) | ks,abs |  |  | 
 |  | 75 V | 200 V | 350 V | 600 V
0.05 | 0.03 | 1.3 | 0.98 | 0.89 | 0.78
0.2 | 0.20 | 3.41 | 1.87 | 1.56 | 1.14
0.25 | 0.14 | 2.46 | 1.33 | 2.05 | -
1 | 0.67 | 6.00 | 3.07 | 2.12 | 1.58
2.2 | 1.25 | 8.80 | 4.12 | 2.80 | 1.94
3 | 1.95 | 11.96 | 5.67 | - | 2.58
4.5 | 2.63 | 14.99 | 6.87 | 4.59 | 3.07
6 | 3.66 | 18.94 | 8.54 | 5.63 | 3.81
7.5 | 4.12 | 19.54 | 8.77 | 5.69 | 3.74
9 | 4.56 | 21.38 | 9.30 | 5.99 | 4.23
10.5 | 5.26 | 22.99 | 9.95 | 6.50 | 4.24
Notes: The calorimeter and chamber were exposed to a quasi-monoenergetic pulsed electron beam of approximately 200 MeV at the CLEAR user facility at CERN [3]. Each electron pulse was formed of a variable number of shorter electron bunches, with an adjustable charge-per-bunch range between 0.001 nC to 1.5 nC. Each bunch had a length of approximately 1 ps. By adjusting the charge-per-bunch and the number of bunches-per-pulse, we investigated charge-per-pulse values ranging from 0.05 nC to 11 nC. The collection time of the ion chamber is of the order of μs, therefore the chamber is sensitive to the total macro pulse and not the individual bunches.
The absolute value of the recombination factor was calculated from the ratio of the absolute dose determined from the calorimeter and the dose calculated from the ion chamber: 𝑘_(𝑠,𝑎𝑏𝑠)=𝐷_(𝑤,𝑐𝑎𝑙)/(𝑀𝑘_𝑝𝑜𝑙 𝑘_𝑇𝑃 〖𝑘_(𝑄,𝑄_0 ) 𝑘_(𝑄,𝑄_0 ) 𝑁〗_(𝐷,𝑤,𝑄_0 ) )
where 𝐷_(𝑤,𝑐𝑎𝑙) is the dose-to-water measured from the calorimeter, M is the charge reading of the chamber, 𝑘_𝑝𝑜𝑙 is the polarity correction factor, 𝑘_𝑇𝑃 is the temperature and pressure correction factor, 𝑘_(𝑄,𝑄_0 )is the beam quality correction factor and 𝑁_(𝐷,𝑤,𝑄_0 ) is the calibration coefficient of the ionisation chamber in the reference beam quality
The dose-to-water conversion factor, used to convert from the measured dose-to-graphite from the calorimeter into the required dose-to-water, was calculated to be 1.0912. The dose-per-pulse to water and the recombination factors determined from absolute measurements are displayed in Table 1.

## Slide 9
ks up to 10 (V = 200V)  collection eff. 10%
ks up to 4 (V = 600V)  collection eff. 25%
ks,abs compared with ks,TVA (two-voltage method)
Available recombination models include Boag’s free-electron fraction models (Boag 1996)
By optimising the free-electron fraction parameter in these equations, we were able to determine a best fit of our data.
All analytical models of Boag and Di Martino show similar predictions of the recombination factor and estimations of the free-electron fraction
Analytical (Boag 1996, Di Martino 2005) and logistic (Petterson 2017) models tested
The logistic model from Petersson shows the best fit to data over the whole dose-per-pulse range, however this model has no physical meaning and simply relies on two fitting constants α and β
Results
Credit: M. McManus

## Slide 10
We carried out ion recombination study for PTW Roos chamber in 200 MeV e-beam
Transfer standard graphite calorimeter has been used as a reference instrument
The Roos chamber exposed to UHDR VHEE suffers from significant collection loses which cannot be described with currently available analytical ion recombination models
Using higher collecting voltage is not solving the problem, but application of ultra-thin gap IC could alleviate the challenge of ion recombination in UHDR VHEEs
Conclusions

## Slide 11: Thank you
anna.subiel@npl.co.uk
This project (18HLT04) has received funding from the EMPIR programme co-financed by the Participating States and from the European Union’s Horizon 2020 research and innovation programme.
http://uhdpulse-empir.eu/
Acknowledgements
Michael McManus
Francesco Romano
Nigel Lee
Hugo Palmans
Wilfrid Farabolini
Antonio Gilardi