## Slide 1: UHDR dosimetry in Antwerp
Verdi Vanreusel1,2,3, A. Gasparini1,2, L. De Freitas Nascimento3, D. Verellen1,2
1 University of Antwerp, Antwerp, Belgium; 2 Iridium Netwerk, Wilrijk, Belgium; 3 SCK CEN, Mol, Belgium;
UHDPulse 2nd
stakeholders meeting
-- 28/01/2023

## Slide 2: Introduction: Antwerp UHDR research
ElectronFlash (ELF)
Detector production + implementation
Dosimetric Framework
In vitro/vivo infrastructures
(Multimodality and image guidance)
Medical Physics
In vitro/vivo models
  Breast cancer
  Skin cancer
3D cultures + organoids
Radiobiology
Clinical translation of FLASH-RT (IOeRT)
Radiation oncology
Slide adapted from A. Gasparini
AReRO

## Slide 3: Introduction: Antwerp accelerator
Slide adapted from A. Gasparini

## Slide 4: Introduction: Antwerp dosimetry
FLASHDiamond
Ionization chambers
Alanine EPR
Optically stimulated luminescence
Point scintillators
2D scintillating system

## Slide 5: Optically stimulated luminescence dosimetry
Semi-flexible sheet
Irradiation in dark
Read-out: CR-reader
Grey Value  Dose
Dark decay  Timing!

## Slide 6
 | OSLhigh | OSLmedium | OSLlow
Particle size (µm) | 3.2 | 1.5 | ±0.2
Phosphor/binder | 12/98 | 5.3/94.7 | 11/99
Thickness (µm) | 33 | 8 | 23
Phosphor coating weight (g/m²) | 4.4 | 0.5 | 3.3
Sensitivity at 60 kVp | 1900 | 180 | 30
Optically stimulated luminescence dosimetry

## Slide 7: OSL dosimetry: Calibration
 | Slope (Conventional) | Slope (UHDR)
OSLhigh | 0.48 ◦ 10-3 | 0.48 ◦ 10-3
OSLmedium | 4.85 ◦ 10-3 | 5.17 ◦ 10-3
OSLlow | 20.53 ◦ 10-3 | 11.18 ◦ 10-3

## Slide 8: OSL dosimetry: Dose limit
OSLmedium and OSLlow  Saturation

## Slide 9: OSL dosimetry: Dose limit

## Slide 10: OSL dosimetry: Linearity with DPP
Linear with dose per pulse up to 10 Gy
 | R²
OSLhigh | 0.9997
OSLmedium | 0.9994
OSLlow | 0.9988

## Slide 11: OSL dosimetry: Stability with PRF
Stability with PRF
 | Maximal deviation (%)
OSLhigh | -0.80
OSLmedium | [-6.63, 3.72]
OSLlow | 1.42

## Slide 12: Point scintillator dosimetry
Scintillating head + lightguide
Read-out: PMT
Intensity  Dose rate
Decay time!

## Slide 13: Point scintillator dosimetry
 | AlO | OrgCl | OrgBr
Material | Al2O3:C,Mg | (C38H34P2)MnCl4 | (C38H34P2)MnBr4
Emission wavelength [nm] | 420 | 525 | 525
Decay time [ms] | 35 | 5 | 0.4
Picture |  |  | 

## Slide 14: Point scintillators: Calibration
 | Slope (Conventional) | Slope (UHDR)
AlO | 51.80 ◦ 10-9 | 66.49 ◦ 10-9
OrgCl | 3.63 ◦ 10-6 | 4.76 ◦ 10-6
OrgBr | 3.93 ◦ 10-6 | 12.84 ◦ 10-6
Vanreusel et al. Physica Medica; 2022

## Slide 15: Point scintillators: Linearity with DPP
Vanreusel et al. Physica Medica; 2022

## Slide 16: Point scintillators: Stability with PRF
Vanreusel et al. Physica Medica; 2022

## Slide 17: 2D scintillator dosimetry
Scintillating sheets
Read-out: camera
Intensity  Dose rate
Decay time!

## Slide 18: 2D scintillator dosimetry
 | AlO | OrgBr | Triazine
Material | Al2O3:C,Mg | (C38H34P2)MnBr4 | C3H3N3
Emission wavelength [nm] | 420 | 525 | ?
Decay time [ms] | 35 | 0.4 | ?
Picture |  |  | 

## Slide 19: 2D scintillator dosimetry
 | AlO | OrgBr | Triazine
Sensitivity (visual) |  |  | 
Gain | 2800 | 100 | 1000
Mean grey value | 6700 | 25000 | 18000
Relative yield | 1 | 104.5 | 7.5
Homogeneity | 6.5% | 5.4% | 5.9%

## Slide 20: 2D scintillators: Linearity with DPP
 | R²
AlO | 0.89
OrgBr | 0.98
Triazine | 0.98

## Slide 21: 2D scintillators: Stability with PRF
Initial increase
Thereafter strong decrease
Not expected to be due to decay time
Further investigation needed!

## Slide 22: 2D scintillators: Dead time camera
5 ms
50 ms
# pulses: 100
PRF: 10 Hz
PL: 4 µs
20 ms
100 ms

## Slide 23: Conclusion
Parameters | OSL | Point scintillators | 2D scintillators
Calibration | OSLhigh: conv = UHDR | UHDR > Conv = saturation! | Dead time camera More investigation needed
Linearity with dose | Good Limited by digitization process | Good No upper limit detected | More investigation needed
Linearity with dose per pulse | Good up to 10 Gy/pulse | Further investigation needed (IT) | Good for all materials
Stability with PRF | Good stability with PRF | Good for materials with short decay time | More investigation needed

## Slide 24: Dangers of UHDR dosimetry

## Slide 25: Made possible by the FLASH team:
The RDA-team (SCK CEN)
The EL(F)ves (University of Antwerp)
The (baby)SITters (S.I.T.)
Corporate Innovation Office (AGFA)